CHANGING OCCLUSION PRESSURE LIMITS ON ALARIS PUMPS

The purpose of this document is to present the instructions, as provided by biomed, for changing the occlusion pressure limit settings on the Alaris Pumps used in IR as of 2-23-11.  When these pumps are used at VCH, or for other specified reasons, they may be set at lower occlusion pressure limit, which leaves the pumps unable, in most instances, to infuse in intra-arterial thrombolysis cases.  The change that we will be making is restoring the pump to the occlusion pressure setting that has been determined appropriate by biomed, and which is used on all general use pumps at VUMC.

First you will need to enter ‘anesthesia mode’ to make any necessary adjustments to the pump settings.

- With pump off, turn it on while also pressing and holding the ‘SYSTEM ON’ and ‘OPTIONS’ buttons simultaneously

- Press the arrow to the right of “PC Unit”

- Press the arrow key on the right-hand side to change ‘Disabled’ to  ‘Enabled’ next to Anesthesia Mode

- Press the arrow key under ‘EXIT’, at the bottom of the screen

Second, you will change the pressure settings:

- Press the arrow key to the left of ‘Pump Module’

- Press the arrow key below ‘PAGE DOWN’

- Press the arrow key to the right of ‘Pressure Mode’

- Press the arrow key to the right of’ Selectable’

- Use the arrow keys next to Up and Down to adjust the pressure settings.  Biomed sets these both at 525mmhg (the maximum setting), which is the setting to which you should restore it for thrombolysis infusion in adult patients.

- Press the arrow key below ‘CONFIRM’

- Press the arrow key below ‘CONFIRM’

- Press the arrow key below ‘EXIT’

At the end of this process, the pump will power off.  If you are going to be using this pump for intra-arterial infusion in an emergency stroke treatment case, your next steps will be:

- Turn the pump on, pressing “SYSTEM ON”

- When asked if this is a new patient, use the arrow key to the right to select ‘Yes’

- When asked, used the arrow key to select  the ‘Critical Care’ profile, if the critical care profile has already been selected, you will select ‘Yes’

- press the arrow key below ‘CONFIRM’

- If you are prompted to enter patient ID, press the arrow key below ‘EXIT’

- Push the ‘Channel Select’ button

- Press the arrow key to the right of ‘Basic Infusion’

- Enter the rate, as instructed by the physician