S7 START UP PROCEDURE

1. Plug unit into electrical outlet

2. Elevate Screen panel to upright position

3. Move mouse to right side of flat working surface

4. Turn unit on by “pushing in” power button located in the far left corner. Button and mouse will be illuminated by blue light.

5. While unit is powering up–check for the following components.

A.Stealth pedal/outside of unit in bottom bucket

B. Touch & go probe

C. Blunt Probe

D. Reference Frame

6. Screen will come up with 2 selections–Select Cranial Application. Will then be highlighted to green along with the Launch button on bottom right of screen.

7. Click Launch

8. Surgeon Profile will come up next. Select surgeon or Standard Profile. Items selected will be highlighted in green.

9. Click on green highlighted arrow at top of screen to advance to next screen.

10. Select procedure screen(tumor, bx, shunt). Will be highlighted in green. Click on green highlighted arrow to advance to next screen.

11. Select Pt. Screen will now be showing.

Note: Pt. Scans are not typically loaded on unit the night before surgery. Images can be loaded by network transfer to unit, CD or by USB.

*Most are network transfer unless that system is down. If so call either MRI 34557/ CT 26528 to have a CD of scan’s “burned off”. Will need to go and retrieve them from the radiology file room on the !st floor.*

12. Connect Network Transfer Cable to right side of the Stealth unit. Connect the other end to the network port in the room labeled “Stealth”.

13.  Login to IMPAX

A) search for and open the study you wish to transmit.

B) right-click on the image and select “Transmit” then

select “Study”from the floating toolbar that displays.

C) Select the destination from the dropdown list.

D) Click “Transmit”

14. The exam(s) selected will then transfer over to the Stealth and populate in the patient list. Click on the patients name,

15. Exam highlighted in green will be showing in row of boxes to the right.

16. If the exam didn’t transfer to the Stealth successfully or only part of the scan transferred: Then repeat Step 13 above, selecting “Series” instead of “Study” from the floating toolbar on section 13-B.

17. Use mouse roller to scroll thru the exams.

*Note- On right side of Screen will see Pt. Name

MR, type of scan/modality, slice count, date of exam, series, description of exam(MRI,CT,MRI/GAD), Spacing, fov.

Exam detail button-gives some more info.

Modify Images- small screen allows user to adjust light/dark, move, zoom, recenter, reorient.

18. Unless Merging Exams- only highlight exam listed on DCP or schedule sheet. Once this is done click the arrow button to proceed to next screen.

19. Screen/set up equipment. Connect or verify connection of surgeon monitor to stealth station.

*typically will be plugged in and solid green line will be on screen. If not line will be broken orange.

20. Click the forward arrow button. The next screen will have a list of Instruments associated with the procedure selected that are typically used.

21. Advance to next screen with the forward button.

22. Scan images will now be showing in standard unit setup. Coronal, sagittal, axial, 3d. Clicking on button in each screen allows user to change the screen to a different screen view for that box only.

23. To the Top Left of scan’s will find 6 adjustment buttons for scan’s. Click on each one to see where display shows up.

Measure button allow’s measurement of tumor’s/cyst etc. Click once to highlight. Click on edge of tumor then other side and measurement will come up.